## Slide 1
Applicable approaches for traceability assurance in GBON stations
Drago Groselj
Online webinar on moving towards SOFF implementation –
Calibration and field verification of the instruments in GBON stations
12th November 2024

## Slide 2: Metrological traceability
Metrological traceability could be described as a direct link between a result of a measurement made in the field and a result obtained by the calibration process in a calibration laboratory.
It ensures that different measurement methods and instruments used in different countries at different times produce reliable, repeatable, reproducible, compatible and comparable measurement results.
When a measurement result is metrologically traceable, it can be confidently linked to the internationally accepted measurement references.
At the top of the metrological traceability chain there is an internationally defined and accepted reference, in most cases the SI, whose technical and organizational infrastructure has been developed and maintained by BIPM.
NMIs are responsible for maintenance of national standards and dissemination of traceability on the national level, either by themselves or by DIs. DIs are experienced institutes operating at the top of the national metrology system.

## Slide 3: Main objective
The main objective of the calibration strategy for traceability assurance is to ensure the proper traceability of measurement and calibration results to the SI, through an unbroken chain of calibrations, each contributing to the measurement uncertainty.
This strategy applies to meteorological measurements for which a traceability chain to the SI is well established (measurements of temperature, atmospheric pressure, humidity, wind speed, precipitation, … ).

## Slide 4: Responsibility
The responsibility for traceability assurance lies with WMO Members, who should ensure all the required calibrations as well as other necessary steps to achieve the objective of the strategy.
It is up to each NMHS to choose the most suitable approach for its traceability assurance, but ensuring the metrological traceability of all measurement results is strongly recommended.

## Slide 5: Traceability assurance scenarios
Fully assured traceability – target, high confidence level in measurements;
Assured traceability (without accreditation) – good confidence level but some risks; improvement recommended;
Partially assured traceability – poor confidence and high risk; improvement required;
Lack of traceability – level of confidence cannot be assessed; urgent need for improvement.

## Slide 6: Fully assured traceability – target, high confidence level in measurements
BIPM / SI Units
NMI / DI
RIC
NMHS / Cal Lab
Met. sensor
Met. sensor
Met. sensor
CIPM MRA
ISO / IEC 17025
ISO / IEC 17025
Cal. Lab.
ISO / IEC 17025
Following preconditions have to be met to achieve this status:
NMHS has a calibration laboratory.
Laboratory personnel are well trained and competent to properly operate laboratory standards and equipment.
Calibration standards and equipment meet the target uncertainties required for calibrations.
Calibration standards and equipment are regularly calibrated and maintained.
Quality management system, including all the calibration procedures, working instructions and forms, is well documented and applied in laboratory work.
Calibration laboratory is accredited according to ISO/IEC 17025.
Calibration laboratory participates in ILC.

## Slide 7: Assured traceability (without accreditation) – good confidence level but some risks; improvement recommended
BIPM / SI Units
NMI / DI
RIC
NMHS / Cal Lab
Met. sensor
Met. sensor
Met. sensor
CIPM MRA
ISO / IEC 17025
Cal. Lab.
ISO / IEC 17025
– NMHS has a calibration laboratory, or at least portable calibration devices covering the whole range of measured meteorological parameters;
– Laboratory personnel are well trained and competent to properly operate calibration standards and equipment;
– Calibration standards and equipment meet the target uncertainties required for calibrations of meteorological instruments;
– Calibration standards and equipment are regularly calibrated and maintained.
In addition, the following are highly recommended:
– Quality management system, including all the calibration procedures, working instructions and forms, should be documented and applied in laboratory work;
– Although not accredited, calibration facilities should follow the requirements of ISO/IEC 17025;
– Participation in the ILC, which will be of great benefit.

## Slide 8: Partially assured traceability – poor confidence and high risk; improvement required
BIPM / SI Units
NMI / DI
RIC
NMHS / Field verification kit
Met. sensor
Met. sensor
Met. sensor
CIPM MRA
ISO / IEC 17025
Cal. Lab.
ISO / IEC 17025
NMHS has no calibration laboratory.
NMHS has field verification kit.
Field verification kit should be acquired, with the required metrological characteristics regarding field instruments and with a calibration certificate issued by accredited calibration laboratory.
The field verification kit should be regularly calibrated by NMI or DI, RIC, of by any accredited calibration laboratory, or by non-accredited laboratory at the worst case.
The field verification kit should be checked before and after field use.
Personnel designated to operate the field verification kit should take part in appropriate training courses organized by RIC, NMI or other relevant institutions.
Technical procedures for operating the field verification kit should be documented.
Field verifications should be performed on a regular time base.
The results of field verification must be documented.

## Slide 9: Lack of traceability – not appropriate way
Lack of metrological traceability leads to a lack of reliability of meteorological measurements, and consequently, highly reduces confidence in the implications of measurement data such as weather forecasts, warnings, and climate analyses.
Ultimately this brings into question the usefulness to meteorological measurements for the global community. So the consequences of untraceable measurement results are severe.
Therefore, measurement traceability is essential and WMO Members are urged to assure traceability of all the measurements under their responsibility.

## Slide 10
WMO Regional Instrument Centres
RA I Alger (Algeria) Cairo (Egypt) Casablanca (Morocco) Nairobi (Kenya) Gaborone (Botswana)
RA II Beijing (China) * Tsukuba (Japan) *
RA IV Bridgetown (Barbados)
RA V Manila (Philippines) Melbourne (Australia) *
RA VI Toulouse (France) Bratislava (Slovakia) * Ljubljana (Slovenia) *
Munich, Hamburg (Germany) *
Ankara (Turkey) *
RA III Buenos Aires (Argentina)
https://community.wmo.int/activity-areas/imop/Regional_Instrument_Centres
* RIC is accredited according to ISO/IEC 17025.

## Slide 11
An RIC shall assist Members of the Region, and possibly of other Regions, in calibrating their national meteorological standards and related environmental monitoring instruments;
An RIC shall participate in, and/or organize, interlaboratory comparisons, and support instrument intercomparisons following relevant WMO recommendations;
According to relevant recommendations on the WMO Quality Management Framework, an RIC shall make a positive contribution to Members regarding the quality of measurements;
An RIC shall advise Members on enquiries regarding instrument performance, maintenance and the availability of relevant guidance materials;
An RIC shall actively participate, or assist, in the organization of workshops on calibration and maintenance of meteorological and related environmental instruments;
…
WMO ToR for RIC – corresponding functions

## Slide 12
RICs calibration capabilities
 |  |  | Quantity |  |  |  |  |  |  | 
Region | Country |  | Temperature | Relative humidity | Atmospheric pressure | Wind | Precipitation | Solar radation | Other (air quality) | Other (electrical)
RA I | Algeria |  |  |  |  |  |  |  |  | 
 | Egypt |  |  |  |  |  |  |  |  | 
 | Kenya |  |  |  |  |  |  |  |  | 
 | Botswana |  | N/A |  |  |  |  |  |  | 
 | Morocco |  |  |  |  |  |  |  |  | 
RA II | China |  |  |  |  |  |  |  |  | 
 | Japan |  |  |  |  |  |  |  |  | 
RA III | Argentina |  |  |  |  |  |  |  |  | 
RA IV | Barbados |  |  |  |  |  |  |  |  | 
RA V | Australia |  |  |  |  |  |  |  |  | 
 | Philippines |  |  |  |  |  |  |  |  | 
RA VI | France |  |  |  |  |  |  |  |  | 
 | Germany |  |  |  |  |  |  |  |  | 
 | Slovakia |  |  |  |  |  |  |  |  | 
 | Slovenia |  |  |  |  |  |  |  |  | 
 | Turkey |  |  |  |  |  |  |  |  | 
 |  |  |  |  | 46 | Calibration capability without ISO/IEC 17025 accreditation |  |  |  | 
 |  |  |  |  | 29 | Calibration capability with ISO/IEC 17025 accreditation |  |  |  | 
https://community.wmo.int/activity-areas/imop/Regional_Instrument_Centres

## Slide 13
NMI, RIC,
Accreddited laboratory
NMHS Calibration laboratory
External traceability
Internal traceability
Ground stations traceability
NMI, RIC,
Accredited Laboratory
Ground stations
Ground stations
Field verification kit
Field verification kit

## Slide 14: Field verification
Field verification offers the ability to check/verify instrument on site.
Suitable also for NMHSs without calibration labs.
Field verification is not calibration.
Verification is usually done at one point against the working standard by placing the working standard as close to the instrument under verification (IUV) as possible.
Verification kit must be periodically recalibrated in calibration laboratory.
Suitable for traditional and AWS instruments.

## Slide 15: Field verification
Stabilization time must be allowed to reach temperature equilibrium between the working standard and the IUV.
Attention must be paid to the proximity of the working standard to the IUV, the temperature gradients, the airflow, the pressure differences and any other factors that could influence the verification results.
The most important disadvantage is that the inspection is usually limited to one point.
The second disadvantage is that if an error is reported, the IUV should be removed and replaced with a calibrated instrument.

## Slide 16: Specification of field verification kit - example
Robust instruments with display
Battery powered with autonomy up to 8h
Barometer with digital display | To measure barometric pressure
Output | Digital display and RS232
Operating temperature range | -10 to 40°C
Power supply | Rechargeable battery pack
Operation time using battery | >8h
Resolution (display) | 0.01hPa
Measurement range | 600 ... 1100 hPa
Linearity | <0.1 hPa
Hysteresis | <0.1 hPa
Temperature dependence | < 0.1 hPa
Total accuracy | <0.2 hPa
Long term stability | < 0.15 hPa/year
Thermometer with digital display | To measure temperature
Output | Digital display
Power supply | Rechargeable battery pack
Operation time using battery | >8h
Measurement range | ​-40 ... +60 °C
​Temperature sensor | ​Pt100 RTD IEC751 1/3 Class B ​
Resolution | 0.1°C
Hygrometer with digital display | To measure relative humidity
Output | Digital display
Power supply | Rechargeable battery pack
Operation time using battery | >8h
Measurement range | ​0 ... 100 % RH
Resolution | 0.1% RH
Accuracy at 20°C | <1%
Long-term stability | <1 %rh / year

## Slide 17
Instrument | Comparison method | Laboratory uncertainty | Field tolerance
Pressure | Within Stevenson scr. | 0.3hPa | 0.5hPa
Temperature | Within Stevenson scr. | 0.1K | 0.3K
Relative humidity | Within Stevenson scr. | 2% | 4%
Wind direction | Compass | 5% | 10%
Example of field tolerances

## Slide 18
Calibration laboratory | Field verification
Large investment in premises, infrastructure and human resources | Initial calibration of instruments followed by one point field verification
Training of calibration laboratory personnel | Moderate investment in transfer standards for verification
Standard operating procedures for calibration | Periodic recalibration of transfer standard – verification kit (NMI, RIC)
Logistics of equipment to laboratory and back to measuring site | Specific training for personnel providing field verification
ISO/IEC 17025 accreditation | To use calibrated spare instruments
Resources needed for ongoing operation and maintenance – traceability of laboratory standards | Lack of guidance on field verification and field tolerances

## Slide 19
References
Guide to Instruments and Methods of Observation, 2018 edition - Volume I: Measurement of Meteorological Variables,, Annex 1.B, Strategy for traceability assurance https://library.wmo.int/doc_num.php?explnum_id=9869,
International vocabulary of metrology – Basic and general concepts and associated terms (VIM)
https://www.bipm.org/documents/20126/2071204/JCGM_200_2012.pdf/f0e1ad45-d337-bbeb-53a6-15fe649d0ff1
Guide on expression of uncertainty
https://www.bipm.org/documents/20126/2071204/JCGM_100_2008_E.pdf/cb0ef43f-baa5-11cf-3f85-4dcd86f77bd6
WMO Knowledge Sharing Portal
https://community.wmo.int/activity-areas/imop/knowledge-sharing-portal
IOM report 119: Guidance on the computation of calibration uncertainties https://library.wmo.int/index.php?lvl=notice_display&id=17152#.YKdmN6FvFaQ

## Slide 20
Thank you
Merci